# ONE fall detection investigation

Real video evidence and the 20 fps question

30 September 2026  •  Technical research report

The current fall rule needs stronger temporal activity reasoning and a measured inference cadence. The evidence does not establish that swapping in a pose detector, or simply sending more frames, will solve the missed falls and false alarms.

Across 128 evaluation clips, the existing ONE path at 5 Hz detected 37 of 63 annotated falls and triggered on 5 of 65 daily-activity clips. At the idealized current 1.4-second interval, it detected 2 of 63 falls. These are offline, cold-start clip results, not estimates of continuous deployment reliability.

## Main evaluation results

| Method and cadence | TP | FN | FP | TN | Precision | Recall | Median latency s |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ONE at 5 Hz | 37 | 26 | 5 | 60 | 88.1% | 58.7% | 1.20 |
| ONE every 1.4 s | 2 | 61 | 0 | 65 | 100.0% | 3.2% | 3.90 |
| Pose boxes at 5 Hz | 30 | 33 | 5 | 60 | 85.7% | 47.6% | 1.40 |
| Pose boxes every 1.4 s | 3 | 60 | 0 | 65 | 100.0% | 4.8% | 3.70 |

128 evaluation clips: 63 Fall and 65 activities of daily living (ADL). TP/FN are detected/missed fall clips; FP/TN are alerted/quiet ADL clips. Latency uses known-onset detected falls: 36 and 2 for ONE; 29 and 3 for pose, respectively. It excludes processing and network delay.

## What the evidence supports

Sampling matters. The browser currently submits at most one JPEG every 1.4 seconds and skips a tick while a request is active. A camera stream at 20 fps does not mean the backend receives or processes 20 frames per second.

Twenty Hz did not improve recall in the paired test. On the same nine evaluation clips, 5 Hz and 20 Hz each detected 2 of 4 falls; false-positive ADL clips increased from 1 of 5 to 2 of 5. The result is a small diagnostic comparison, not a general ranking of frame rates.

Pose remains an open research option. At 5 Hz, pose boxes detected 30 of 63 falls, versus 37 for ONE, with five false-positive ADL clips for each. The same tracker and box rule were used; a trained temporal skeleton classifier was not tested.

Recommended next step: compare time-based temporal logic and activity models under realistic capture-to-alert timing, including continuous pre-fall history and long non-fall recordings. Preserve a separate evaluation set before choosing thresholds.

# Dataset and evaluation protocol

Accuracy was evaluated on public GMDCSA-24 RGB videos of consenting actors staging falls and daily activities at home. Scenes include low light, seated and chair falls, partial views, lying down, bending, exercise and bed activities. [1, 2]

| Partition | Subject folders | Fall | ADL | Clips | Samples at 5 Hz |
| --- | --- | --- | --- | --- | --- |
| Development | 1 | 16 | 16 | 32 | 1,007 |
| Evaluation | 2 to 4 | 63 | 65 | 128 | 5,432 |
| Total | 1 to 4 | 79 | 81 | 160 | 6,439 |

All 160 video hashes match the pinned revision. They contain 34,172 native frames and about 21.5 minutes of material. Both models completed all 6,439 samples with no missing or duplicate frame IDs. The manifest yields 79 Fall and 81 ADL clips, reversing the paper’s reported totals; the manifest governs scoring.

## Partitioning and prior exposure

Methods were fixed before the initial evaluation. No weights were trained or fine-tuned, and no held-out threshold tuning was performed. The expanded set includes the earlier 32-clip pilot; results beyond the pilot are reported separately.

The split is by subject folder. The paper describes four actors in its main text but three in its ethics statement; identities were not independently verified. Treat this as folder-separated evaluation, without a guarantee of person independence. [1]

## Replay and scoring

Each clip resets detector-associated tracking and fall state. Frames are resized to a maximum dimension of 640 pixels and encoded at JPEG quality 58, approximating the browser canvas quality of 0.58. The real low-light preprocessing, detector, stability tracker and fall rule are replayed offline using source timestamps. This is not a full HTTP or app deployment test.

A fall clip is positive when the rule emits an alert at or after the source fall onset minus 0.2 seconds; ADL is positive for any alert. These are clip-level event labels, not person-localized fall ground truth. Latency is the first qualifying alert time minus the annotation time. One fall clip, s4_fall_15, lacks onset timing: it remains in classification counts and is omitted from latency.

## Cold start changes the interpretation

Three observations need a minimum span of 2.8 seconds at the 1.4-second interval, 0.4 seconds at 5 Hz and 0.1 seconds at 20 Hz. Of 62 evaluation falls with onset timing, 41 begin before 2.8 seconds; neither detector path caught any of these at the 1.4-second interval. These results do not estimate continuous-monitoring performance.

# Expanded results and diagnostic comparisons

The following diagnostics retain the same clip labels and scoring rule. They isolate potential failure mechanisms but are not production-ready alternatives.

| Method and cadence | TP | FN | FP | TN | Precision | Recall | Median latency s |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ONE at 5 Hz | 37 | 26 | 5 | 60 | 88.1% | 58.7% | 1.20 |
| One identity at 5 Hz | 40 | 23 | 5 | 60 | 88.9% | 63.5% | 1.20 |
| Confidence gate removed at 5 Hz | 41 | 22 | 5 | 60 | 89.1% | 65.1% | 1.20 |
| Pose boxes at 5 Hz | 30 | 33 | 5 | 60 | 85.7% | 47.6% | 1.40 |
| Pose with one identity at 5 Hz | 33 | 30 | 7 | 58 | 82.5% | 52.4% | 1.40 |

Identity diagnostic: force one ID onto the highest-confidence person prediction. Confidence diagnostic: keep detector outputs above the 0.20 detection floor, but bypass the fall rule’s 0.55 gate. Neither is a deployment recommendation.

## What the ablations show

For ONE, forcing one identity recovered three additional falls; removing the confidence gate recovered four, with unchanged false-positive clip counts. For pose, forcing one identity recovered three falls but added two false-positive clips. Association and confidence gating contribute to misses, but most missed falls remained. These diagnostics do not establish safety on multiple people or long non-fall recordings.

The pilot had 8 development and 24 evaluation clips. ONE detected 9 of 12 falls at 5 Hz with one false-positive ADL clip. The expanded results provide a less favorable and more useful basis for the next experiment.

## Variation within the expanded set

On the 104 evaluation clips beyond the pilot, ONE detected 28 of 51 falls and alerted on 4 of 53 ADL clips. Pose boxes detected 22 of 51 and alerted on 3 of 53. These additional examples account for the lower full-set recall.

For clips marked night, ONE detected 5 of 11 falls and pose boxes 4 of 11. By day, they detected 32 of 52 and 26 of 52, respectively. Scene and activity also vary, so this comparison does not isolate lighting as a cause.

## Metrics that should not be overread

Each base detector path has five false-positive ADL clips containing seven alerts. Counting each false alert but only one detection per fall clip gives 84.1% precision for ONE and 81.1% for pose, versus 88.1% and 85.7% by clip. The artificial fall prevalence means neither estimates alert quality in an occupied home. False-positive clips do not estimate false alarms per camera-day.

Counts of frames containing a person prediction are coverage diagnostics, not person-detection recall. There are no verified person boxes for that claim, and the pose model sometimes labels hanging clothes as another person. Reported event latency excludes processing, network and UI delay, and only includes detected falls.

# Faster sampling and runtime feasibility

Twelve clips were selected before 20 Hz inference: three development and nine evaluation. The evaluation subset has four falls and five ADL clips, with 1,287 samples. Lower cadences use paired subsamples of the same cached detector outputs.

| Method and cadence | TP | FN | FP | TN | Precision | Recall | Median latency s |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ONE every 1.4 s | 0 | 4 | 0 | 5 | — | 0.0% | — |
| ONE at 5 Hz | 2 | 2 | 1 | 4 | 66.7% | 50.0% | 3.25 |
| ONE at 20 Hz | 2 | 2 | 2 | 3 | 50.0% | 50.0% | 0.75 |
| ONE at 20 Hz with 9 hits | 1 | 3 | 1 | 4 | 50.0% | 25.0% | 1.15 |

Nine-clip subset only. One fall changes recall by 25 percentage points. Latency is conditional on detection and excludes processing delay. A dash means undefined.

Keeping three hits while increasing cadence shortens the rule’s minimum evidence span. The separate nine-hit diagnostic at 20 Hz restores a 0.4-second minimum span, but does not make every aspect of the temporal logic duration-invariant. It detected one fall and raised one false-positive clip. Higher cadence alone did not resolve the activity ambiguity.

All nine evaluation videos have native gaps below 50 ms. The three development videos have a few larger gaps, where resampling may hold a frame. Twenty-nine full-set videos are about 15 fps and cannot supply 20 independent observations per second.

On the full set, shifting the 1.4-second sampling grid across seven phase offsets produced only 0 to 2 detected falls for ONE and 0 to 3 for pose, with no false-positive ADL clips.

## Measured CPU cost

| Runtime path | Per frame time | Approximate reciprocal |
| --- | --- | --- |
| Current path on four blank smoke frames | 2.01 to 2.18 s | About 0.5 fps |
| Exact-vocabulary embedding cache | 267.9 ms | 3.7 fps |
| Static prompt configuration | 147.6 ms | 6.8 fps |

Shared Linux CPU, AMD EPYC 9V74, two Torch threads. The last two times are means from one predeclared development frame, two warmups and ten repetitions. Smoke timings are a different workload and are not directly comparable.

Static prompts matched parsed detections exactly on five predeclared development frames and cut mean per-frame time by about 45% versus the cached comparison path. This is limited equivalence evidence, not a universal guarantee. Timing includes JPEG decode, low-light processing and inference; it excludes capture, HTTP, tracking, persistence, display, file read, resize and JPEG encoding.

Serial 20 Hz inference needs roughly 50 ms per frame; this CPU path averaged 148 ms. Benchmark the complete request path on deployment hardware before changing cadence. No Mac, GPU or end-to-end throughput result is claimed.

# Observed failure mechanisms and next experiments

## Low confidence after the fall

In pilot clip s3_fall_16, the detector produced eight upright, one transition and nine low-posture boxes above its 0.20 floor. Only the upright boxes survived the fall rule’s 0.55 confidence gate. Removing that gate recovered the fall. The next test should assess calibrated confidence and low-light temporal evidence while measuring additional false alarms.

## Seated and partial views

Clip s3_fall_13 began in a low seated posture and stayed low through the fall; lowering confidence did not recover it. The code can fall back to the previous center when no upright reference exists, so an upright observation is not an absolute prerequisite. Its box geometry still lacks a useful transition here. The partial chair fall s4_fall_10 remained missed after either the identity or confidence diagnostic.

## Ordinary activity that resembles a fall

The push-up clip s4_adl_05 triggered both ONE and the pose-box path at 4.4 seconds in the 5 Hz pilot replay. Better person detection does not remove this activity-specificity problem. A falling body and deliberate exercise can share a low box and downward movement.

## Experiments to run before choosing an approach

1  Measure the live timing path. Record capture, upload, inference, tracking and alert timestamps, skipped requests, queue age and p50/p95 latency. Compare 5, 10 and 20 Hz targets on the actual deployment hardware using newest-frame processing rather than an accumulating queue.

2  Make evidence duration explicit. Compare time-based stabilization and fall-confirmation windows, with continuous pre-roll, against the unchanged rule. Isolate cadence from the amount of temporal evidence instead of reusing an unqualified frame count.

3  Compare temporal representations. Evaluate RGB motion, tracked box trajectories and a trained temporal skeleton classifier on the same partitions. Include seated falls, occlusion, low light, bed transitions and exercise. Pose should earn its role through this comparison.

4  Use a fresh external evaluation. Preserve untouched people and environments before tuning. Add long ADL streams and report false alarms per camera-hour, missed falls, onset latency and performance by scenario. Use repeated runs or multiple seeds for trained models.

None of these findings establishes medical or emergency-response reliability. The investigation is evidence for the next engineering experiment, not a production fix or a safety validation.

# Reproducibility and limits

## Pinned sources and licenses

ONE backend source: 9c05becdc3af59b15092075c1eb1dc90e2ab9e70. Frontend inspected: 84e21f90f53213c79775168ee0491019ad1f1400. The replay uses the fixed eleven-label default profile; a live home with registered objects may request a different vocabulary. [3]

GMDCSA-24 v2.1 is pinned to 5abac7693229900cf80f722e878fbb119211fc1c. Its Zenodo record licenses the dataset under CC BY 4.0 and notes revised CSV activity timings. Cite that v2.1 record and the paper. The repository separately carries an MIT notice, copyright 2024 Ekram Alam; preserve that notice with reproduced repository material. [1, 2, 5]

## Weights and software

yolov8s-worldv2.pt  SHA256

9b2c17ab6124a913e9b3a5c170617920d91b0f01111a8479da69f00e2cf27792

yolo11n-pose.pt  SHA256

869e83fcdffdc7371fa4e34cd8e51c838cc729571d1635e5141e3075e9319dc0

CLIP ViT B 32  SHA256

40d365715913c9da98579312b702a82c18be219cc2a73407c4526f58eba950af

Both YOLO checkpoints came from official Ultralytics v8.4.0 assets. CLIP fork revision: a13192f8cb767260d7dfd98c843b0716593169e7. Runtime: Python 3.12.14, Torch 2.14.0+cpu, torchvision 0.29.0+cpu, Ultralytics 8.4.168, OpenCV 4.14.0 and NumPy 2.5.2. The companion lockfile records exact packages. [4]

Detector settings: 384 by 640 input, 0.20 detection floor and 100 maximum detections. Tracker: three hits, IoU at least 0.20. Fall rule: person confidence at least 0.55; box width and height each at least 8% of frame dimensions; upright height/width at least 1.20; low at most 0.95; vertical-center drop at least 0.07 of frame height; three low confirmations within four seconds.

## Limits that affect decisions

Short staged clips, coarse onset annotations, uncertain actor mapping, overlapping pilot exposure, one missing onset and no verified person-level fall annotations constrain the conclusions. One CSV lists s2_adl_04 as 70 seconds; its decoded duration is about 7.84 seconds. Native timestamps and decoded durations are used where metadata conflicts. This investigation makes no production-code changes, merge or deployment.

## Sources and companion evidence

[1] Alam et al 2024 GMDCSA 24 paper

[2] Ekram Alam 2024 GMDCSA-24 v2.1 dataset and CC BY 4.0 license

[3] Pinned ONE backend source

[4] Ultralytics YOLO World documentation

[5] Pinned GitHub repository and separate MIT notice

Paper DOI: 10.1016/j.dib.2024.110892. Companion evidence includes per-clip manifests and hashes, source snapshots, inference outputs, scores, protocol declarations, runtime benchmarks and the pilot diagnostic contact sheet and video.